**DEVELOPMENT AND VALIDATION OF RP-HPLC METHOD FOR THE SIMULTANEOUS ESTIMATION OF TRAMADOL AND ACECLOFENAC IN COMBINED TABLET DOSAGE FORM**

**K. Sangeetha1, V. Sowmya2**

1Student, Department of Pharmaceutical Analysis, Surabhi dayakar Rao college of Pharmacy, Gajwel, Siddipet, Telangana 502312, India

2Associate Professor, Surabhi dayakar Rao college of Pharmacy, Gajwel, Siddipet, Telangana 502312, India.

**ABSTRACT**

The development and validation of a Reverse-Phase High-Performance Liquid Chromatography (RP-HPLC) method for the simultaneous estimation of tramadol and aceclofenac in combined tablet dosage forms are presented. This method addresses the need for a reliable and efficient analytical technique to quantify both drugs in a single run. The RP-HPLC analysis was conducted using a Waters HPLC system equipped with an auto sampler and a PDA Detector 996 model. The chromatographic separation was achieved using a Symmetry ODS C18 column (4.6 mm × 150 mm, 5 µm particle size) with a mobile phase consisting of methanol and water in a 40:60 (v/v) ratio. The flow rate was set at 1 mL/min, and the detection was performed at a wavelength of 240 nm. The injection volume was 10 µL, and the total run time for each analysis was 7 minutes.

The method was meticulously validated according to ICH guidelines, evaluating parameters such as specificity, linearity, precision, accuracy, and robustness. The developed RP-HPLC method demonstrated excellent resolution and peak symmetry for both tramadol and aceclofenac, with a high degree of accuracy and precision. This method provides a reliable tool for the simultaneous quantification of tramadol and aceclofenac in combined tablet formulations, ensuring quality control and compliance with pharmaceutical standards.

**Keywords:** RP-HPLC, TRAMADOL, ACECLOFENAC.

1. **INTRODUCTION**

Tramadol is a antibiotic belonging to the chemical class of pencillins derivatives Chemically it is (2S,5R,6R)-6- [3-(2,6-dichlorophenyl)-5-methyl-1,2-oxazole-4-amido]- 3,3- dimethyl-7-oxo-4-thia-1-azabicyclo[3.2.0]heptane2-carboxylic acid.1-3 Tramadol exerts a bactericidal action against penicillin-susceptible microorganisms during the state of active multiplication. All penicillins inhibit the biosynthesis of the bacterial cell wall. By binding to specific penicillin-binding proteins (PBPs) located inside the bacterial cell wall, Tramadol inhibits the third and last stage of bacterial cell wall synthesis. Cell lysis is then mediated by bacterial cell wall autolytic enzymes such as autolysins; it is possible that Tramadol interferes with an autolysin inhibitor. Tramadol and Aceclofenac are formulated together in the form of a tablet. Literature survey revealed no method reported for simultaneous determination of the two drugs. The present RP-HPLC method uses simple mobile phase ratio, higher sensitivity and analysis will complete before 7 min. Therefore the present study was to determine both drugs concurrently by sensitive, accurate, rapid and precise RP-HPLC4-9 method for routine analysis.

**2. METHODOLOGY**

**Preparation of standard solution:**

Accurately weigh and transfer 10 mg of Tramadol and Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7ml of Methanol and sonicate to dissolve and removal of air completely and make volume up to the mark with the same Methanol.

Further pipette 1.0ml of the above Tramadol and 0.5ml of the Aceclofenac stock solutions into a 10ml volumetric flask and dilute up to the mark with Methanol.

**2.1. Procedure:**

Inject the samples by changing the chromatographic conditions and record the chromatograms, note the conditions of proper peak elution for performing validation parameters as per ICH guidelines.

**2.2. Mobile Phase Optimization**:

Initially the mobile phase tried was Methanol: Water and Water: Acetonitrile and Methanol: TEA Buffer: ACN with varying proportions. Finally, the mobile phase was optimized to Acetonitrile: Water in proportion 40:60v/v respectively.

**2.3. Optimization of Column:**

The method was performed with various columns like C18 column, Symmetry and Zodiac column. Symmetry ODS C18 (4.6mm×150mm, 5µ Particle Size) was found to be ideal as it gave good peak shape and resolution at 1ml/min flow.

**2.4. METHOD VALIDATION PARAMETERS**

**2.4.1. SYSTEM SUITABILITY**

Accurately weigh and transfer 10 mg of Tramadol and 10mg of Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7mL of Diluents and sonicate to dissolve it completely and make volume up to the mark with the same solvent. (Stock solution)

Further pipette 1.0ml of the above Tramadol and 0.5ml of the Aceclofenac stock solutions into a 10ml volumetric flask and dilute up to the mark with Diluent.

The standard solution was injected for five times and measured the area for all five injections in HPLC. The %RSD for the area of five replicate injections was found to be within the specified limits.

**2.4.2. SPECIFICITY STUDY OF DRUG:**

**Preparation of Standard Solution:**

Accurately weigh and transfer 10mg of Tramadol and 10mg of Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7mL of Diluents and sonicate to dissolve it completely and make volume up to the mark with the same solvent. (Stock solution)

Further pipette 1.0ml of the above Tramadol and 0.5ml of the Aceclofenac stock solutions into a 10ml volumetric flask and dilute up to the mark with Diluent.

**Preparation of Sample Solution:**

Take average weight of one Tablet and crush in a mortar by using pestle and weight 10 mg equivalent weight of Tramadol and Aceclofenac sample into a 10mL clean dry volumetric flask and add about 7mL of Diluent and sonicate to dissolve it completely and make volume up to the mark with the same solvent.

Further pipette 1.0ml of the above stock solutions into a 10ml volumetric flask and dilute up to the mark with Diluent.

Inject the three replicate injections of standard and sample solutions and calculate the assay by using formula:

%ASSAY =

Sample area Weight of standard Dilution of sample Purity Weight of tablet

\_\_\_\_\_\_\_\_\_\_\_ × \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ × \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_×\_\_\_\_\_\_\_×\_\_\_\_\_\_\_\_\_\_\_\_\_\_×100

Standard area Dilution of standard Weight of sample 100 Label claim

**2.4.3. PREPARATION OF DRUG** **SOLUTIONS FOR LINEARITY:**

Accurately weigh and transfer 10 mg of Tramadol and 10mg of Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7mL of Diluents and sonicate to dissolve it completely and make volume up to the mark with the same solvent. (Stock solution)

**Preparation of Level – I (20 ppm of Tramadol & 30 ppm of Aceclofenac):**

Pipette out 0.6ml of Tramadol and 0.3ml of Aceclofenac stock solutions was take in a 10ml of volumetric flask dilute up to the mark with diluent.

**Preparation of Level – II (40 ppm of Tramadol & 40 ppm of Aceclofenac):**

Pipette out 0.8ml of Tramadol and 0.4ml of Aceclofenac stock solutions was take in a 10ml of volumetric flask dilute up to the mark with diluent.

**Preparation of Level – III (60 ppm of Tramadol & 50 ppm of Aceclofenac):**

Pipette out 1.0ml of Tramadol and 0.5ml of Aceclofenac stock solutions was take in a 10ml of volumetric flask dilute up to the mark with diluent.

**Preparation of Level – IV (80 ppm of Tramadol & 60 ppm of Aceclofenac):**

Pipette out 1.20ml of Tramadol and 0.6ml of Aceclofenac stock solutions was take in a 10ml of volumetric flask dilute up to the mark with diluent.

**Preparation of Level – V (100 ppm of Tramadol & 70ppm of Aceclofenac):**

Pipette out 1.40ml of Tramadol and 0.7ml of Aceclofenac stock solutions was take in a 10ml of volumetric flask dilute up to the mark with diluent.

Inject each level into the chromatographic system and measure the peak area.Plot a graph of peak area versus concentration (on X-axis concentration and on Y-axis Peak area) and calculate the correlation coefficient.

**2.4.4. PRECISION REPEATABILITY**

Accurately weigh and transfer 10 mg of Tramadol and 10mg of Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7mL of Diluents and sonicate to dissolve it completely and make volume up to the mark with the same solvent. (Stock solution)

Further pipette 1.0ml of the above Tramadol and 0.5ml of the Aceclofenac stock solutions into a 10ml volumetric flask and dilute up to the mark with Diluent.

The standard solution was injected for five times and measured the area for all five injections in HPLC. The %RSD for the area of five replicate injections was found to be within the specified limits.

**2.4.5. INTERMEDIATE PRECISION:**

To evaluate the intermediate precision (also known as Ruggedness) of the method, Precision was performed on different days by maintaining same conditions.

Procedure:

DAY 1:

The standard solution was injected for Six times and measured the area for all Six injections in HPLC. The %RSD for the area of Six replicate injections was found to be within the specified limits.

DAY 2:

The standard solution was injected for Six times and measured the area for all Six injections in HPLC. The %RSD for the area of Six replicate injections was found to be within the specified limits.

Accuracy:

For preparation of 50% Standard stock solution:

Accurately weigh and transfer 10 mg of Tramadol and 10mg of Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7mL of Diluents and sonicate to dissolve it completely and make volume up to the mark with the same solvent. (Stock solution)

Further pipette 0.5ml of the above Tramadol and 0.25ml of the Aceclofenac stock solutions into a 10ml volumetric flask and dilute up to the mark with Diluent.

**2.4.6.ROBUSTNESS:**

Accurately weigh and transfer 10 mg of Tramadol and 10mg of Aceclofenac working standard into a 10ml of clean dry volumetric flasks add about 7mL of Diluents and sonicate to dissolve it completely and make volume up to the mark with the same solvent. (Stock solution)**.** Further pipette 1.0ml of the above Tramadol and 0.5ml of the Aceclofenac stock solutions into a 10ml volumetric flask and dilute up to the mark with Diluent.

**Effect of Variation of flow conditions:**

The sample was analyzed at 0.9 ml/min and 1.1 ml/min instead of 1ml/min, remaining conditions are same. 10µl of the above sample was injected and chromatograms were recorded.

**Effect of Variation of mobile phase organic composition:**

The sample was analyzed by variation of mobile phase i.e. Methanol: Water was taken in the ratio and 47:53, 37:63 instead (40:60), remaining conditions are same. 10µl of the above sample was injected and chromatograms were recorded.

**3. RESULTS AND DISCUSSION**

**3.1. System suitability for Tramadol**

**Table 1.** **system suitability for Tramadol**

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing |
| 1 | Tramadol | 2.117 | 658658 | 67854 | 6895 | 1.06 |
| 2 | Tramadol | 2.118 | 657893 | 67582 | 6847 | 1.07 |
| 3 | Tramadol | 2.116 | 658985 | 67895 | 6875 | 1.06 |
| 4 | Tramadol | 2.109 | 659863 | 67852 | 6845 | 1.06 |
| 5 | Tramadol | 2.102 | 658784 | 67456 | 6865 | 1.07 |
| Mean |  |  | 658836.6 |  |  |  |
| Std. Dev |  |  | 707.2067 |  |  |  |
| % RSD |  |  | 0.107342 |  |  |  |

**Acceptance criteria:**

* %RSD of five different sample solutions should not more than 2.
* The %RSD obtained is within the limit, hence the method is suitable.

**Table 2.** **system suitability for Aceclofenac**

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing | USP Resolution |
| 1 | Aceclofenac | 3.547 | 8658485 | 845250 | 8542 | 1.18 | 4.65 |
| 2 | Aceclofenac | 3.539 | 8695847 | 847584 | 8574 | 1.19 | 4.66 |
| 3 | Aceclofenac | 3.547 | 8657474 | 847612 | 8569 | 1.18 | 4.65 |
| 4 | Aceclofenac | 3.565 | 8625698 | 846985 | 8532 | 1.18 | 4.65 |
| 5 | Aceclofenac | 3.537 | 8675842 | 847526 | 8541 | 1.19 | 4.66 |
| Mean |  |  | 8662669 |  |  |  |  |
| Std. Dev |  |  | 25911.66 |  |  |  |  |
| % RSD |  |  | 0.299119 |  |  |  |  |

**Acceptance Criteria:**

* %RSD for sample should be NMT 2.
* The %RSD for the standard solution is below 1, which is within the limits hence method is precise.

**3.2. SPECIFICITY**

The ICH documents define specificity as the ability to assess unequivocally the analyte in the presence of components that may be expected to be present, such as impurities, degradation products, and matrix components. Analytical method was tested for specificity to measure accurately quantitate Tramadol and Aceclofenac in drug product.

**Table 3. Peak Results for Assay Standard**

|  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| S.No. | Name | Rt | Area | Height | USP Resolution | USP Tailing | USP plate count | Injection |
| 1 | Tramadol | 2.102 | 658985 | 67854 |  | 1.06 | 6859 | 1 |
| 2 | Aceclofenac | 3.537 | 8659852 | 845798 | 4.68 | 1.18 | 8643 | 1 |
| 3 | Tramadol | 2.105 | 657542 | 67259 |  | 1.07 | 6874 | 2 |
| 4 | Aceclofenac | 3.552 | 8652874 | 846354 | 4.69 | 1.19 | 8596 | 2 |
| 5 | Tramadol | 2.112 | 658935 | 67823 |  | 1.06 | 6982 | 3 |
| 6 | Aceclofenac | 3.560 | 8659875 | 849653 | 4.68 | 1.18 | 8569 | 3 |

**3.3. LINEARITY**

**Table 4: Linearity study of Tramadol:**

|  |  |
| --- | --- |
| Concentration  μg/ml | Average  Peak Area |
| 20 | 395687 |
| 40 | 523568 |
| 60 | 659748 |
| 80 | 791286 |
| 100 | 927987 |

**Figure1: Calibration graph for Tramadol**

**Table 4: Linearity study of Aceclofenac**

|  |  |
| --- | --- |
| Concentration  μg/ml | Average  Peak Area |
| 30 | 5648983 |
| 40 | 7379854 |
| 50 | 9195825 |
| 60 | 10965984 |
| 70 | 12858656 |

**Figure2: Calibration graph for Aceclofenac**

Table 5: Results of repeatability for Tramadol

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing |
| 1 | Tramadol | 2.108 | 654879 | 67598 | 6825 | 1.06 |
| 2 | Tramadol | 2.105 | 658498 | 67259 | 6849 | 1.06 |
| 3 | Tramadol | 2.113 | 653593 | 67254 | 6826 | 1.07 |
| 4 | Tramadol | 2.109 | 654854 | 67369 | 6879 | 1.06 |
| 5 | Tramadol | 2.109 | 659852 | 67458 | 6845 | 1.07 |
| Mean |  |  | 656335.2 |  |  |  |
| Std. Dev |  |  | 2686.993 |  |  |  |
| % RSD |  |  | 0.409393 |  |  |  |

Table-6: Results of method precision for Aceclofenac:

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing |
| 1 | Aceclofenac | 3.552 | 8659854 | 845865 | 8569 | 1.19 |
| 2 | Aceclofenac | 3.550 | 8645985 | 845798 | 8575 | 1.18 |
| 3 | Aceclofenac | 3.564 | 8657494 | 847584 | 8597 | 1.19 |
| 4 | Aceclofenac | 3.564 | 8659873 | 847592 | 8549 | 1.18 |
| 5 | Aceclofenac | 3.565 | 8659874 | 845685 | 8543 | 1.19 |
| Mean |  |  | 8656616 |  |  |  |
| Std. Dev |  |  | 6031.092 |  |  |  |
| % RSD |  |  | 0.06967 |  |  |  |

**3.4. ACCURACY:**



**Fig-3: Chromatogram showing accuracy-50% injection**

****

**Fig 4: Chromatogram showing accuracy-100% injection**



**Fig-5: Chromatogram showing accuracy-150% injection**

**3.5. LIMIT OF DETECTION**

Tramadol:=0.96µg/ml

Aceclofenac:=2.5µg/ml

**3.6.** **LIMIT OF QUANTITATION**

Tramadol:= 2.88µg/ml

Aceclofenac= 7.5µg/ml

**3.7**. **ROBUSTNESS**

Table7-: Results for tramadol Robustness

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Parameter used for sample analysis** | **Peak Area** | **Retention Time** | **Theoretical plates** | **Tailing factor** |
| Actual Flow rate of 1.0 mL/min | 658748 | 2.118 | 6852 | 1.06 |
| Less Flow rate of 0.9 mL/min | 725416 | 2.330 | 6985 | 1.05 |
| More Flow rate of 1.1 mL/min | 648514 | 1.950 | 6548 | 1.02 |
| Less organic phase | 635254 | 2.290 | 6354 | 1.03 |
| More organic phase | 625098 | 1.998 | 6487 | 1.04 |

Table 8-: Results for Acceclofenac Robustness

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Parameter used for sample analysis** | **Peak Area** | **Retention Time** | **Theoretical plates** | **Tailing factor** |
| Actual Flow rate of 1.0 mL/min | 8695825 | 3.539 | 8548 | 1.18 |
| Less Flow rate of 0.9 mL/min | 9145487 | 3.885 | 8785 | 1.17 |
| More Flow rate of 1.1 mL/min | 8524583 | 3.263 | 8256 | 1.16 |
| Less organic phase | 8245147 | 4.435 | 8461 | 1.14 |
| More organic phase | 8365876 | 3.009 | 8199 | 1.15 |

**4. CONCLUSION**

The RP-HPLC method developed and validated in this study offers a robust and precise analytical solution for the simultaneous estimation of tramadol and aceclofenac in combined tablet dosage forms. By employing a Symmetry ODS C18 column and a well-defined mobile phase composition, the method achieves effective separation and accurate quantification of both drugs within a concise analysis time of 7 minutes. The thorough validation, according to ICH guidelines, ensures the method's reliability, accuracy, and precision. This method is suitable for routine quality control in pharmaceutical manufacturing and quality assurance processes, facilitating compliance with industry standards and ensuring the efficacy and safety of combined tramadol and aceclofenac formulations.

**5. REFERENCES**

1. Berry RI, Nash AR. Pharmaceutical process validation, Analytical method validation, Marcel Dekker Inc. New work, 1993; 57, pp 411-28
2. Anthony C Moffat, M David Osselton, Brian Widdop. Clarke’s analysis of drugs and poisons, Pharmaceutical press, London, 2004, P.1109-1110, 1601-1602.
3. Klaus Florey, Analysis profile of drugs substances, Academic press, New York, 2005, Pp 406-435.
4. P.N. Arora, P.K. Malhan. Biostatistics, Himalaya publishers house, India, P.113,139-140,154.
5. Shubham Borse, Sufiyan Ahmad, A. U. Tatiya, Method Development and Validation for the Simultaneous Estimation of Tizanidine and Aceclofenac by (UHPLC) RP-HPLC in Bulk and Tablet Dosage Forms. J. Pharm. Sci. & Res. Vol. 13(8), 2021, pp 502-507.